9480 Carroll Park Dr., San Diego, CA 92121-5201

### Service Note

Phone (858) 657-9800   Fax: (858) 657-9818 Toll Free 1800 324-5190

E-mail:  api-customerservice@teledyne.com     http://www.teledyne-api.com

- 10-019

- 2 December 2010

T Series Windows CE Update

- Purpose:

To give instruction on installing new Window CE software in a T Series instrument.

- TOOLS:

Computer with Ethernet communication.

USB standard keyboard

- PARTS:

Teledyne API Instrument Upgrade Utility

Latest revision of Firmware from Teledyne API Customer Service

The electronics used in T-API analyzers are sensitive to Electrostatic Discharge (ESD). When working on any T-API device, please ensure that you are properly grounded prior to handling or touching any electronic circuitry in the analyzers! For more information on how to protect sensitive components from ESD during handling, please contact T-API customer service and ask for the ESD Service note number 03-022A.

- PROCEDURE:

- The OS Image Update must be completed through the Ethernet connection. Connect the instrument to the computer via a patch cable or a hub by network cable.

- Locate the Instrument Upgrade Utility install file, Setup.msi, and double click. Follow the onscreen instructions. Note: Must have the latest windows update of .net framework installed.

- Return to your desktop, or where you have downloaded the new Windows CE image file. Open the new file. Rename the ***.bin file to NK.bin.

- Copy down the IP address of the instrument. The IP address can be found by pressing SETUP on the instrument’s display, MORE, COM and INET.  Enter 9-2-9 password and ENTER. Press SET> till INST IP: is displayed and note this number ***.***.***.***.

- Click the Start menu at the lower left-hand of the screen: All Programs, Teledyne API and Instrument Upgrade Utility.

- The Teledyne API Instrument Upgrade Utility will open on the screen. Click Next, Instrument OS Image and Next. Instrument Firmware Image page will appear. Click the square “…”. This will allow you to search for the NK.bin file from your desktop, or your saved file. Select the NK.bin file. The Select Target Instrument page will now appear. Type in the IP Address of the instrument, which you copied from step 4, and select Next. A confirmation page will show up, click Next and Next again. The Performing Update screen will appear showing the status of the update. When finished, click Finish, and at the next screen prompt click OK. Turn the instrument off by the front panel switch. The touch screen will now need to be calibrated continue to the next step for calibrating the display.

Touchscreen Calibration

- Connect a keyboard to one of the USB ports on the front panel. Turn the instrument back on and hold down the Left CTRL and Left SHIFT keys on the keyboard when the screen turns white. Continue holding till the calibration window appears.

- Press and release the center crosshair. A timer function is enabled, and this only allows a 15 second time to start the calibration process. If the timer expires the instrument will continue into its normal operation.

- A new crosshair will be displayed in the upper left-hand corner of the screen. Press on the center of the crosshair with one finger. The Display will prompt you to remove you finger once the point is configured. Continue this process till all four corner calibrations are completed.

- Two buttons will appear in the lower left corner, CANCEL and ACCEPT. Press ACCEPT and cycle the power on the front panel. If you press CANCEL, the instrument will boot, but the calibration won’t be accepted.