9480 Carroll Park Dr., San Diego, CA 92121-5201

### Service Note

Phone (858) 657-9800   Fax: (858) 657-9818 Toll Free 1800 324-5190

E-mail:  api-customerservice@teledyne.com     http://www.teledyne-api.com

- 10-020

- 2 December 2010

T Series Firmware Update

- Purpose:

To give instruction on installing new operating firmware in a T Series instrument.

- TOOLS:

Computer with Ethernet communication or USB flashdrive

USB keyboard

- PARTS:

Teledyne API Instrument Upgrade Utility

Latest revision of Firmware from Teledyne API Customer Service

The electronics used in T-API analyzers are sensitive to Electrostatic Discharge (ESD). When working on any T-API device, please ensure that you are properly grounded prior to handling or touching any electronic circuitry in the analyzers! For more information on how to protect sensitive components from ESD during handling, please contact T-API customer service and ask for the ESD Service note number 03-022A.

- PROCEDURE:

  - The firmware update can be completed through Ethernet using the Instrument Upgrade Utility, or it can be done with a USB flash drive and the front panel USB ports.

Ethernet

- Locate the Instrument Upgrade Utility install file, Setup.msi, and double click. Follow the onscreen instructions. (Note: Must have the latest windows update of .net framework installed.)

- Copy down the IP address of the instrument. The IP address can be found by pressing SETUP on the instrument’s display, MORE, COM and INET.  Enter 9-2-9 password and ENTER. Press SET> till INST IP: is displayed and note this number ***.***.***.***.

- Click the Start menu at the lower left-hand of the screen: All Programs, Teledyne API and Instrument Upgrade Utility.

- The Teledyne API Instrument Upgrade Utility will open on the screen. Click Next, Instrument Application Firmware and Next. Instrument Firmware Image page will appear. Click the square “…”. This will allow you to search for the new firmware file from your desktop, or your saved file. Select the appropriate .exe file. The Select Target Instrument page will now appear. Type in the IP Address of the instrument, which you copied from step 4, and select Next. A confirmation page will show up, click Next and Next again. The Performing Update screen will appear showing the status of the update. When finished, click Finish, and at the next screen prompt click OK. Power cycle the instrument via the from panel switch.

USB Flashdrive

- Open the new firmware file and Rename the ***.exe file to “update.exe”.

- Insert a USB flashdrive into the computer and copy the “update.exe” file to the root directory of the flashdrive. Remove the flash drive from the computer and install it into the analyzer’s front USB port. NOTE: The analyzer dose not check firmware and it will install any file named “update.exe” so double-check to insure you are installing the correct firmware.

- A screen will appear on the display of the analyzer confirming the update of firmware. If this is correct press YES.

- Enter 1-2-3-4, ENTER. The instrument will load the new software, and then prompt to remove the USB flashdrive.

- The instrument will restart itself.